## Slide 1
0
1
2
15
16
MIKUMARI port numbers
TDC 0-31 ch
TDC 32-63 ch
16
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: Scaler reset
NIM IN2: Trigger (veto) in
NIM OUT1: Heartbeat signal
NIM OUT2: TCP connection
MIKUMARI Clock Primary
0-16: Primary (TX)
LED1: PLL locked
LED2: 0
LED3: Heartbeat frame state
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Not in Use
DIP3: Not in Use
DIP4: Not in Use
＊前のバージョンにあったMIKUMARIの状態を記録するTDCは廃止

## Slide 2
0
1
2
15
16
MIKUMARI port numbers
TDC 0-31 ch
TDC 32-63 ch
16
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: TDC 64 ch
NIM IN2: Trigger (veto) in
NIM OUT1: Heartbeat signal or Trigger out (from LACCP)
NIM OUT2: TCP connection
MIKUMARI Clock Hub
0-15: Primary (TX)
16: Secondary (RX)
LED1: PLL locked
LED2: MIKUMARI (16) link up
LED3: Ready for DAQ
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Trigger out
DIP3: Not in Use
DIP4: Not in Use
＊前のバージョンにあったMIKUMARIの状態を記録するTDCは廃止

## Slide 3
MIKUMARI port numbers
AMANEQ StrLrTdc
LED1: PLL locked
LED2: MIKUMARI (0) link up
LED3: Ready for DAQ
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Trigger out
DIP3: Stand alone mode
DIP4: Not in Use
0
TDC 0-31 ch
TDC 32-63 ch
0
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: TDC 128 ch
NIM IN2: Trigger in
NIM OUT1: Heartbeat signal or Trigger out (from LACCP)
NIM OUT2: TCP connection
0: Secondary (RX)
TDC 64-95 ch
TDC 96-127 ch

## Slide 4
今回の運用
MikuClockPrim, MikuClockHub
ParingをOFF。Trailing dataが返ってくる。
AMANEQ StrLrTdc, CIRASAME StrTdc
ParingをON。TOTが入ったleading dataだけ返ってくる。

## Slide 5
0
1
2
15

## Slide 7
FPGA
Power
supply
SFP1
2
Mezzanine card
Differential signal to LVCMOS
32 differential signal input
Mezzanine
base connectors
DC +20-35V
AC/DC adaptor jack
2
AMANEQ module
LVDS
DIP Switch
SFP2
Coaxial
Coaxial
SFP3
GTX
Logic
translate
NIM
Logic
translate
Jitter
cleaner
SD
RAM
4
32
32
32
32
4
Mini-mezzanine card

## Slide 8
Local bus
FMP
Local bus
controller
TCP
RBCP
SiTCP/XG
Inside
FPGA
SEM
XADC
SDS
CECE62002
Controller
DAQ
functions
CECE
62002
SFP
MIKUMARI
LACCP
Heartbeat
Unit
MIKUMARI
Util
Flash
memory
MMCM
MMCM
Oscillator
100MHz
Local clock
domain
Global clock
domain
PCS/PMA
SFP
CDC

## Slide 9
16
15
Secondary
MIKUMARI
Primary MIKUMARIs
Status bits
CbtLaneUp
MikumariUp
LaccpUp
If RegIndex is 15,
CbtTapValueIn
CbtTapValueOut
CbtBitSlipIn
PartnerIpAddr
NumLinks = 17
Offset information
of secondary
HbcOffset
LocalFineOffset
LaccpFineOffset
Control bits
RstOverMiku

## Slide 10
MIKUMARI
link
MIKUMARI
link
TX
RX
PLL
User circuit
Reference
system clock
Serial clock
Oscillator
RX
TX
PLL
Modulated clock
signal
User circuit
Transmission line
Recovered
system clock
Primary side
Secondary side
Source pulse
Reproduced pulse at destination
Reproduced pulse edge is correctly synchronized with the source pulse.
~205 ns (3m fiber)

## Slide 11
Heartbeat
Counter
(16-bit)
Heartbeat pulse
DAQ function
Heartbeat frame
Delimiter
(has 24-bit
frame number)
DAQ data
LACCP
MIKUKARI
PLL
Heartbeat
Counter
(16-bit)
Heartbeat pulse
DAQ function
Heartbeat frame
Delimiter
(has 24-bit
frame number)
DAQ data
LACCP
MIKUKARI
Clock sync.
PLL
Frequency sync.
Reference system clock
0xfffe
0xffff
0x0
0x1
0x2
Heartbeat
Heartbeat
counter
Primary side
Recovered system clock
0xfffe
0xffff
0x0
0x1
0x2
Heartbeat
Heartbeat
counter
Secondary side
Local fine offset
Frame state

## Slide 12
Lower Main
input
Upper Main
input
Upper mezzanine slot
Lower mezzanine slot
Streaming
TDC
MIKUMARI
LACCP
Heartbeat
Unit
MMCM
Scaler
TCP
RBCP
SiTCP
NIMIN-1
System
clock
Local oscillator clock
NIMIN-2
Trigger

## Slide 13
Heartbeat delimiter
Front merger
(8-to-1)
Front merger
(8-to-1)
Front merger
(8-to-1)
Front merger
(8-to-1)
Back merger
(16-to-1)
Merging (Vital) block
Buffer
SiTCP
ODP block
Heartbeat
Unit
Delimiter
generator
LACCP
MIKUMARI
・
・
・
・
・
・
: FIFO

## Slide 14
DAQ functions
Timing from TDC unit + 16-bit fine-scale timestamp
16-bit fine-scale timestamp
(8ns, 524us)
Heartbeat
TDC data
Heartbeat delimiter
(24-bit coarse-scale timestamp)
TDC data

## Slide 15
Timing Unit
Timing Unit
Path
merger
Delay
Delimiter
inserter
TOT
filter
Leading edge
Trailing edge
Data path merging
Hit signal
In a channel
Paring
Trigger gate
2us
16-bit fine-scale timestamp
Delimiter

## Slide 16
Back
merger
Front merger
Front merger
Front merger
Front merger
SiTCP
(SiTCP-XG)
ODP block
: FIFO
(32-to-1)
(4-to-1)

## Slide 17
Front merger
Front merger
Front merger
Back merger
Vital block
Front merger
Link Buffer
SiTCP
Output
throttling
Input throttling
type2
Incoming
buffer
Stop the TDC data
Programmable full flag
Empty flag
Programmable full flag
Empty flag
Input throttling type2 is working
HBF
throttling
Delimiter
Inserter
・・・
・・・
・・・
・・・
・・・
・・・
・・・
・・・
・・・
・・・
・・・
・・・