## Slide 1
0
1
2
15
16
MIKUMARI port numbers
TDC 0-31 ch
TDC 32-63 ch
16
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: Scaler reset
NIM IN2: Trigger (veto) in
NIM OUT1: Heartbeat signal
NIM OUT2: TCP connection
MIKUMARI Clock Primary
0-16: Primary (TX)
LED1: PLL locked
LED2: 0
LED3: Heartbeat frame state
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Not in Use
DIP3: Not in Use
DIP4: Not in Use
＊前のバージョンにあったMIKUMARIの状態を記録するTDCは廃止

## Slide 2
0
1
2
15
16
MIKUMARI port numbers
TDC 0-31 ch
TDC 32-63 ch
16
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: TDC 64 ch
NIM IN2: Trigger (veto) in
NIM OUT1: Heartbeat signal or Trigger out (from LACCP)
NIM OUT2: TCP connection
MIKUMARI Clock Hub
0-15: Primary (TX)
16: Secondary (RX)
LED1: PLL locked
LED2: MIKUMARI (16) link up
LED3: Ready for DAQ
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Trigger out
DIP3: Not in Use
DIP4: Not in Use
＊前のバージョンにあったMIKUMARIの状態を記録するTDCは廃止

## Slide 3
0
MIKUMARI port numbers
TDC 0-31 ch
TDC 32-63 ch
0
*GN-2006-1
*GN-2006-4
SFP1
SFP2
Data link
Not in use
NIM
IN
NIM
OUT
These power supply connectors are connected.
Recommend to put a cover on unused side for your safety.
NIM IN1: TDC 128 ch
NIM IN2: Trigger in
NIM OUT1: Heartbeat signal or Trigger out (from LACCP)
NIM OUT2: TCP connection
AMANEQ StrLrTdc
0: Secondary (RX)
LED1: PLL locked
LED2: MIKUMARI (0) link up
LED3: Ready for DAQ
LED4: DAQ is running
DIP1: SiTCP IP setting
DIP2: Trigger out
DIP3: Stand alone mode
DIP4: Not in Use
TDC 64-95 ch
TDC 96-127 ch

## Slide 4
今回の運用
MikuClockPrim, MikuClockHub
ParingをOFF。Trailing dataが返ってくる。
AMANEQ StrLrTdc, CIRASAME StrTdc
ParingをON。TOTが入ったleading dataだけ返ってくる。

## Slide 5
0
1
2
15

## Slide 7
FPGA
Power
supply
SFP1
2
Mezzanine card
Differential signal to LVCMOS
32 differential signal input
Mezzanine
base connectors
DC +20-35V
AC/DC adaptor jack
2
AMANEQ module
LVDS
DIP Switch
SFP2
Coaxial
Coaxial
SFP3
GTX
Logic
translate
NIM
Logic
translate
Jitter
cleaner
SD
RAM
4
32
32
32
32
4
Mini-mezzanine card

## Slide 8
Local bus
FMP
Local bus
controller
TCP
RBCP
SiTCP/XG
Inside
FPGA
SEM
XADC
SDS
CECE62002
Controller
DAQ
functions
CECE
62002
SFP
MIKUMARI
LACCP
Heartbeat
Unit
MIKUMARI
Util
Flash
memory
MMCM
MMCM
Oscillator
100MHz
Local clock
domain
Global clock
domain
PCS/PMA
SFP
CDC